## Slide 1: Analysis of LCP Mosquito Performance on Mitegen InSitu Plates
Shanshan Yu
Jan08, 2018

## Slide 2
Left drop missing count
(aggregation of 5 plates)
1 2 3 4 5 6 7 8 9 10 11 12 A 2 1 0 1 1 4 2 1 2 1 2 1 B 0 0 0 0 0 0 0 0 0 0 0 0 C 0 1 0 0 0 0 0 0 0 0 0 0 D 0 0 0 0 0 0 0 0 0 0 0 0 E 0 0 0 0 0 0 0 0 1 0 0 0 F 0 0 0 0 0 0 0 0 0 1 0 0 G 0 0 0 1 0 0 0 0 0 0 0 0 H 0 0 0 0 0 0 0 2 1 0 0 0
Right drop missing count
(aggregation of 5 plates)
1 2 3 4 5 6 7 8 9 10 11 12 A 0 0 0 0 0 0 2 1 2 1 1 2 B 0 0 0 0 0 1 1 1 0 0 2 1 C 0 0 0 0 0 1 1 1 0 0 0 0 D 0 0 0 0 0 0 0 0 0 0 0 1 E 0 0 0 0 0 0 0 0 0 0 1 1 F 0 0 0 0 0 0 0 0 0 0 1 1 G 0 0 1 0 0 0 0 0 0 0 0 2 H 0 0 0 0 0 0 0 0 0 0 0 1
Plate # | Drop volume (ul) | Left drop | Right drop | Drop Ingredient
1 | 1.5 | 1 | 1 | PEG4K, water
2 | 1 | 1 | 1 | PEG4K, 10mg/ml lysozyme
3 | 1 | 1 | 1 | PEG6K, 10mg/ml lysozyme
4 | 1 | 1 | 1 | PEG4K, 10mg/ml lysozyme
5 | 1 | 1 | 1 | PEG6K, 10mg/ml lysozyme
Plate 1
Red frame are the wells with one or two drops missing
Missing Drop Count
Plates set up between Dec 19-23, 2017
Two sitting drop per well using a 3-drop sitting protocol
Left drop at position 1, right drop at position 3

## Slide 3
The views of two wells with missing drops under a microscope
Plate 1,
red frame are the wells C2 and E12
I suspect that the missing drops are because the distance between mosquito needles and walls of wells are too close. It could also because needles touched the walls during pipetting.
Missing Drop Close View

## Slide 4
Plate 1
The view of four wells with both sitting drops
The distances between two drops look good, but both of them are still close to the walls inside wells. Could the two drops be adjusted to be a little further away from walls?
Other Drop Close View

## Slide 5
x axis
y axis
1️⃣⃣⃣
2️⃣⃣⃣⃣
3️⃣⃣⃣⃣⃣
x-axis
y-axis
The level was On bench, under the humidity chamber
1️⃣⃣⃣
x-axis
y-axis
1, 2, 3 are three positions that are measured by a level
LCP mosquito bench surface is horizontal

## Slide 6
x axis
y axis
1️⃣⃣⃣
2️⃣⃣⃣⃣
3️⃣⃣⃣⃣⃣
x axis
y axis
2️⃣⃣⃣⃣
The level was on top of the humidity chamber frame, right hand side to the LCP syringe
The bubble sit at the bottom, indicating the y-axis not horizontal
1, 2, 3 are three positions that are measured by a level
LCP mosquito plate deck titles towards the back

## Slide 7
x axis
y axis
1️⃣⃣⃣
2️⃣⃣⃣⃣
3️⃣⃣⃣⃣⃣
3️⃣⃣⃣⃣⃣
The level was on top of the humidity chamber frame, left hand side to the pipetting head
x axis
y axis
The bubble sit at the bottom, indicating the y-axis not horizontal
1, 2, 3 are three positions that are measured by a level
LCP mosquito plate deck titles towards the back